## Slide 1: Safety Engineering

## Slide 2: Safety
Safety is a property of a system that reflects the system’s ability to operate, normally or abnormally, without danger of causing human injury or death and without damage to the system’s environment.
It is important to consider software safety as most devices whose failure is critical now incorporate software-based control systems.

## Slide 3: Software in safety-critical systems
The system may be software-controlled so that the decisions made by the software and subsequent actions are safety-critical. Therefore, the software behaviour is directly related to the overall safety of the system.
Software is extensively used for checking and monitoring other safety-critical components in a system. For example, all aircraft engine components are monitored by software looking for early indications of component failure. This software is safety-critical because, if it fails, other components may fail and cause an accident.

## Slide 4: Safety and reliability
Safety and reliability are related but distinct
  In general, reliability and availability are necessary but not sufficient conditions for system safety
Reliability is concerned with conformance to a given specification and delivery of service
Safety is concerned with ensuring system cannot cause damage irrespective of whether or not it conforms to its specification.
  System reliability is essential for safety but is not enough
  Reliable systems can be unsafe

## Slide 5: Safety specification
The goal of safety requirements engineering is to identify protection requirements that ensure that system failures do not cause injury or death or environmental damage.
Safety requirements may be ‘shall not’ requirements i.e. they define situations and events that should never occur.
Functional safety requirements define:
  Checking and recovery features that should be included in a system
  Features that provide protection against system failures and external attacks

## Slide 6: Hazard-driven analysis
Hazard identification
Hazard assessment
Hazard analysis
Safety requirements specification

## Slide 7: Hazard identification
Identify the hazards that may threaten the system.
Hazard identification may be based on different types of hazard:
  Physical hazards
  Electrical hazards
  Biological hazards
  Service failure hazards
  Etc.

## Slide 8: Insulin pump risks
Insulin overdose (service failure).
Insulin underdose (service failure).
Power failure due to exhausted battery (electrical).
Electrical interference with other medical equipment (electrical).
Poor sensor and actuator contact (physical).
Parts of machine break off in body (physical).
Infection caused by introduction of machine (biological).
Allergic reaction to materials or insulin (biological).

## Slide 9: Hazard assessment
The process is concerned with understanding the likelihood that a risk will arise and the potential consequences if an accident or incident should occur.
Risks may be categorised as:
  Intolerable. Must never arise or result in an accident
  As low as reasonably practical(ALARP). Must minimise the possibility of risk given cost and schedule constraints
  Acceptable. The consequences of the risk are acceptable and no extra costs should be incurred to reduce hazard probability

## Slide 10: The risk triangle

## Slide 11: Social acceptability of risk
The acceptability of a risk is determined by human, social and political considerations.
In most societies, the boundaries between the regions are pushed upwards with time i.e. society is less willing to accept risk
  For example, the costs of cleaning up pollution may be less than the costs of preventing it but this may not be socially acceptable.
Risk assessment is subjective
  Risks are identified as probable, unlikely, etc. This depends on who is making the assessment.

## Slide 12: Hazard assessment
Estimate the risk probability and the risk severity.
It is not normally possible to do this precisely so relative values are used such as ‘unlikely’, ‘rare’, ‘very high’, etc.
The aim must be to exclude risks that are likely to arise or that have high severity.

## Slide 13: Risk classification for the insulin pump
Identified hazard | Hazard probability | Accident severity | Estimated risk | Acceptability
1.Insulin overdose computation | Medium | High | High | Intolerable
2. Insulin underdose computation | Medium | Low | Low | Acceptable
3. Failure of hardware monitoring system | Medium | Medium | Low | ALARP
4. Power failure | High | Low | Low | Acceptable
5. Machine incorrectly fitted | High | High | High | Intolerable
6. Machine breaks in patient | Low | High | Medium | ALARP
7. Machine causes infection | Medium | Medium | Medium | ALARP
8. Electrical interference | Low | High | Medium | ALARP
9. Allergic reaction | Low | Low | Low | Acceptable

## Slide 14: Hazard analysis
Concerned with discovering the root causes of risks in a particular system.
Techniques have been mostly derived from safety-critical systems and can be
  Inductive, bottom-up techniques. Start with a proposed system failure and assess the hazards that could arise from that failure;
  Deductive, top-down techniques. Start with a hazard and deduce what the causes of this could be.

## Slide 15: Fault-tree analysis
A deductive top-down technique.
Put the risk or hazard at the root of the tree and identify the system states that could lead to that hazard.
Where appropriate, link these with ‘and’ or ‘or’ conditions.
A goal should be to minimise the number of single causes of system failure.

## Slide 16: An example of a software fault tree

## Slide 17: Fault tree analysis
Three possible conditions that can lead to delivery of incorrect dose of insulin
  Incorrect measurement of blood sugar level
  Failure of delivery system
  Dose delivered at wrong time
By analysis of the fault tree, root causes of these hazards related to software are:
  Algorithm error
  Arithmetic error

## Slide 18: Risk reduction
The aim of this process is to identify dependability requirements that specify how the risks should be managed and ensure that accidents/incidents do not arise.
Risk reduction strategies
  Hazard avoidance;
  Hazard detection and removal;
  Damage limitation.

## Slide 19: Strategy use
Normally, in critical systems, a mix of risk reduction strategies are used.
In a chemical plant control system, the system will include sensors to detect and correct excess pressure in the reactor.
However, it will also include an independent protection system that opens a relief valve if dangerously high pressure is detected.

## Slide 20: Insulin pump - software risks
Arithmetic error
  A computation causes the value of a variable to overflow or underflow;
  Maybe include an exception handler for each type of arithmetic error.
Algorithmic error
  Compare dose to be delivered with previous dose or safe maximum doses. Reduce dose if too high.

## Slide 21: Examples of safety requirements
SR1: The system shall not deliver a single dose of insulin that is greater than a specified maximum dose for a system user.
SR2: The system shall not deliver a daily cumulative dose of insulin that is greater than a specified maximum daily dose for a system user.
SR3: The system shall include a hardware diagnostic facility that shall be executed at least four times per hour.
SR4: The system shall include an exception handler for all of the exceptions that are identified in Table 3.
SR5: The audible alarm shall be sounded when any hardware or software anomaly is discovered and a diagnostic message, as defined in Table 4, shall be displayed.
SR6: In the event of an alarm, insulin delivery shall be suspended until the user has reset the system and cleared the alarm.